## 表紙～目次
[11 empty] | 技術資料番号 |  | ：
[5 empty] | 文書名 |  | ： | 単体テスト仕様書
[5 empty] | 名称 |  | ： | DCMTK置き換え対応
[5 empty] | 形名 |  | ： | XA-161
| 改訂番号 |  |  | 年 月 日 |  |  |  |  | 承 認 |  | 検 印 |  | 担 当
| 0 |  |  | 2015/04/20 |  |  |  |  | FPT-PhuocMT |  | FPT-PhuocMT |  | FPT-ThuanPV3
[6 empty] | 承 認 |  |  | 検 印 | 担 当
[6 empty] | FPT 15.04.21 PhuocMT |  |  | FPT 15.04.21 PhuocMT | FPT 15.04.20 ThuanPV3
履 歴
改訂番号 |  |  | 内 容 | [8 empty] | 作成日、担当
0 |  |  | 初版 | [8 empty] | 2015/04/20 |  | ThuanPV3
目 次
1 概要
1-1 目的
1-2 検査方法
1-3 関連文書
2 テスト項目
2-1 モジュールテスト
2-1-1 CMntLogDlg.cpp
2-1-2 CMntTransDlg.cpp
2-1-3 DcmtkWrapAssociation.cpp
2-1-4 DcmtkWrapCommand.cpp

## 1概要
1 概要
1-1 目的
DBの再構築対応で追加/修正したモジュールの単体での動作の確認
1-2 検査方法
|  | ① | VisualStudio2013のデバッガを用いて、テスト項目の判定基準を満たすこと
を確認する。
|  | ② | 2-1モージュールテストは、検査モジュールの頭でBreakして検査条件を設定
した後に、実行を再開してアウトプットおよび、モジュール内のローカル変数
の設定を確認する。
|  | ③ | CodeCoverageを用いて全ての分岐を通ることを確認する。
|  | ④ | 次の項目に関してソフトウェアコードチェックを行う
ソフトウェアコードが、ソフトウェアユニットの詳細設計で文書化されている
インターフェースと矛盾していないか
ソフトウェアコードが、コーディング標準に従っているか
ソフトウェアコードが、実装で以下問題がないこと
・型変換によるデータ落ちがないこと
・未定義エリアの参照がないこと
ソフトウェアユニットの詳細設計で文書化されている処理内容に対し、
不足または不要な処理が無いか
ソフトウェアユニットの詳細設計で文書化されているモジュール以外に
修正を行っていないか
(モジュールごとではなく、ファイル単位で実施する。)
1-3 関連文書
・プログラム設計書_DCMTK置き換え対応_改訂01(SFT-105269).docx

## モジュール
2 テスト項目
2-1 モジュールテスト
2-1-1 MntLogDlg.cpp
\proper\mainte\MntLogDlg.cpp
No | 関数名 | 関数を呼び出す手順 | 検査条件 | 判定基準 | 備考
1 | void OnInitDialog () | - This function is called to initialize Log dialog | - None | - Path to the configuration file is created.
2 | BOOL CheckDICOMLogExist() | - This function is called to check the existing of DICOM log files | -First selected item position is not NULL - this position != NULL - SelItem = List->GetNextSelectedItem(pos); - List->GetItemText(SelItem, mntLOG_TYPE) == _T("DICOM") | - Return TRUE
3 | int ProcessDICOMLogFile(TCHAR * DrivePath) | - This function is called to processing DICOM log file | - First selected item position is not NULL - Enough free space on disk | - Total space are calculated - Free space are calculated
4 |  |  | - SelItem = List->GetNextSelectedItem(pos); - List->GetItemText(SelItem, mntLOG_TYPE) == _T("DICOM") | - Source path and destination path are created
5 |  |  | - Size of file > total space | - Messagebox are displayed for warning
6 |  |  | - Size of file < total space - Size of file > free space | - Loop for insert next disk - function CopyDICOMLogToMedia is called to copy DICOM log file to media
7 | int CopyDICOMLogToMedia(TCHAR *szSrcFile, TCHAR *szDestFile) | - This function is called to remove Patient Name from DICOM logs and copy to media | - Old file and new file are created - Cannot open old or new file | - return -1
8 |  |  | - Read line and not EOF - Patient Name tag is found in line | - Patient Name in line is replaced by string '*' with the same length - Line is written to new file
9 |  |  | - Copy old file to media failed | - Messagebox is displayed - Delete old file and rename new file to old file
10 | void OnBtnCopyToMedia1() | - This function is called to copy log file to media | - DICOM log files are selected - Click Yes on messagebox | - Function ProcessDICOMLogFile is called
11 |  |  | - Return of function is -1 | - Return -1
12 | int AddListItemFolder(TCHAR *Path) | - This function is called to display the information in the log files in the folder to the list. | - File name contain string'.log' - - File name contain string'.Sck' - - File name contain string'.log.' | - Get some information and add file name to list view
13 | int CopyToMedia(TCHAR * DrivePath) | - This function is called to copy the files that have been selected in the list to the media. | - m_bRemovePatientInfo = TRUE or m_bAllowOutputLog = FALSE - List->GetItemText(SelItem, mntLOG_TYPE) == _T("DICOM") | - continue the loop
2-1-2 MntTransDlg.cpp
\proper\mainte\MntTransDlg.cpp
No | 関数名 | 関数を呼び出す手順 | 検査条件 | 判定基準 | 備考
1 | CMntTransDlg(CWnd * pParent) | - Constructor method | - None | - Path to configuration file (dcmtk.cfg) is created - Code for create path to dicom folder is commented out
2 | BOOL OnInitDialog() | - This function is called to initializing process | - None | - Code for disable two buttons m_ButtonS1 and m_ButtonM1 are commmented out
3 | void TransSCheck1() | - This function is called when Transfer Syntax (Single) "Implicit Little" button is pressed | - m_ButtonS2.GetCheck() == 0 - m_ButtonS3.GetCheck() == 0 - m_ButtonS4.GetCheck() == 0 | - Force m_ButtonS1 button turn to ON - return
4 | void TransSCheck2() | - This function is called when Transfer Syntax (Single) "Explicit Little" button is pressed | - m_ButtonS1.GetCheck() == 0 - m_ButtonS3.GetCheck() == 0 - m_ButtonS4.GetCheck() == 0 | - Force m_ButtonS2 button turn to ON - return
5 | void TransSCheck3() | - This function is called when Transfer Syntax (Single) "RLE Lossless" button is pressed | - m_ButtonS1.GetCheck() == 0 - m_ButtonS2.GetCheck() == 0 - m_ButtonS4.GetCheck() == 0 | - Force m_ButtonS3 button turn to ON - return
6 | void TransSCheck4() | - This function is called when Transfer Syntax (Single) "JPEG Baseline" button is pressed | - m_ButtonS1.GetCheck() == 0 - m_ButtonS2.GetCheck() == 0 - m_ButtonS3.GetCheck() == 0 | - Force m_ButtonS4 button turn to ON - return
7 | void TransMCheck1() | - This function is called when Transfer Syntax (Multi) "Implicit Little" button is pressed | - m_ButtonM2.GetCheck() == 0 - m_ButtonM3.GetCheck() == 0 - m_ButtonM4.GetCheck() == 0 | - Force m_ButtonM1 button turn to ON - return
8 | void TransMCheck2() | - This function is called when Transfer Syntax (Multi) "Explicit Little" button is pressed | - m_ButtonM1.GetCheck() == 0 - m_ButtonM3.GetCheck() == 0 - m_ButtonM4.GetCheck() == 0 | - Force m_ButtonM2 button turn to ON - return
9 | void TransMCheck3() | - This function is called when Transfer Syntax (Multi) "RLE Lossless" button is pressed | - m_ButtonM1.GetCheck() == 0 - m_ButtonM2.GetCheck() == 0 - m_ButtonM4.GetCheck() == 0 | - Force m_ButtonM3 button turn to ON - return
10 | void TransMCheck4() | - This function is called when Transfer Syntax (Multi) "JPEG Baseline" button is pressed | - m_ButtonM1.GetCheck() == 0 - m_ButtonM2.GetCheck() == 0 - m_ButtonM3.GetCheck() == 0 | - Force m_ButtonM4 button turn to ON - return
11 | void GetTransferSyntax() | - This function is called to get transfersyntax from configuration file | - None | - Code for set ON to Implicit Little button is commented out
12 |  |  | - Get number of single Transfersyntax - nIndex < nTransferSyntaxNum - dwRtn > 0 - _tcscmp(szData, MNT_TRANS_IMPLICITLITTLE) == 0 | - mb_SingleFlg_i[0] = 1; - mb_SingleFlg_n[0] = 1;
13 |  |  | - Get number of single Transfersyntax - nIndex < nTransferSyntaxNum - dwRtn > 0 - _tcscmp(szData,MNT_TRANS_EXPLICITLITTLE) == 0 | - mb_SingleFlg_i[1] = 1; - mb_SingleFlg_n[1] = 1;
14 |  |  | - Get number of single Transfersyntax - nIndex < nTransferSyntaxNum - dwRtn > 0 - _tcscmp(szData,MNT_TRANS_RLELOSSLESS) == 0 | - mb_SingleFlg_i[2] = 1; - mb_SingleFlg_n[2] = 1;
15 |  |  | - Get number of single Transfersyntax - nIndex < nTransferSyntaxNum - dwRtn > 0 - _tcscmp(szData, MNT_TRANS_JPEG) == 0 | - mb_SingleFlg_i[3] = 1; - mb_SingleFlg_n[3] = 1;
16 |  |  | - Get number of multi Transfersyntax - nIndex < nTransferSyntaxNum - dwRtn > 0 - _tcscmp(szData, MNT_TRANS_IMPLICITLITTLE) == 0 | - mb_MultiFlg_i[0] = 1; - mb_MultiFlg_n[0] = 1;
17 |  |  | - Get number of multi Transfersyntax - nIndex < nTransferSyntaxNum - dwRtn > 0 - _tcscmp(szData, MNT_TRANS_EXPLICITLITTLE) == 0 | - mb_MultiFlg_i[1] = 1; - mb_MultiFlg_n[1] = 1;
18 |  |  | - Get number of multi Transfersyntax - nIndex < nTransferSyntaxNum - dwRtn > 0 - _tcscmp(szData, MNT_TRANS_RLELOSSLESS) == 0 | - mb_MultiFlg_i[2] = 1; - mb_MultiFlg_n[2] = 1;
19 |  |  | - Get number of multi Transfersyntax - nIndex < nTransferSyntaxNum - dwRtn > 0 - _tcscmp(szData, MNT_TRANS_JPEG) == 0 | - mb_MultiFlg_i[3] = 1; - mb_MultiFlg_n[3] = 1;
20 | void SetTransferSyntax() | - This function is called to set transfersyntax to configuration file | - nIndex < 4 | - Write NULL value to remove un-used fields from configuration file
21 | int GetTransferSyntaxNum(TCHAR* szAppName, TCHAR* szConfigFile) | - This function is called to get number of transfer syntax items | - dwRtn > 0 | - Get value of transfer syntax in configuration file - nIndex++
2-1-3 DcmtkWrapAssociation.cpp
\common\DcmtkWrap\DcmtkWrapAssociation.cpp
No | 関数名 | 関数を呼び出す手順 | 検査条件 | 判定基準 | 備考
1 | int dcwInitializeAscParam(void ** pNetwork, void ** pAscParam, const wchar_t * pszLocalAE, const wchar_t * pszLocalHost, const wchar_t * pszRemoteAE, const wchar_t * pszRemoteHost, const int nRemotePortNum, const int nTimeout) | - This function is called to initialize association parameters | - Initialize network failed | - Write string "Failed to initialize network." to log file
2 |  |  | - Create association falied | - Write string "Failed to create association parameters." to log file
3 |  |  | - Set AE titles failed | - Write string "Failed to set AE titles." to log file
4 |  |  | - Set host name failed | - Write string "Failed to set host name." to log file
5 |  |  | - Set transport layer failed | - Write string "Failed to set transport layer" to log file
6 | int dcwOpenAssociation(void * network, void * assocParam, void ** pAssociation, const wchar_t * profile) | - This function is called to opent association | - Set association parameters failed | - Write string "Failed to set association parameters." to log file
7 |  |  | - None | - Write log for paramters
8 |  |  | - CondRet == DUL_REQUESTASSOCIATIONFAILED | - Write string "Failed to set host name" to log file
9 |  |  | - Request association failed | - Write string "Failed to association reques.t" to log file
10 |  |  | - None | - Write log for paramters
11 |  |  | - ASC_countAcceptedPresentationContexts(param) == 0 | - Write string "All presentation contexts is rejected." to log file
12 | int dcwCloseAssociation(void** pAscociation) | - This function is called to close association | - None | - Write string "Sent association release Request." to log file
13 |  |  | - cond.bad() = TRUE | - Write log for release association error
14 | int dcwAbortAssociation(void** pAssociation) | - This function is called to abort association | - None | - Write string "Sent association abort Request." to log file (info log)
15 | int dcwInitializeNetwork(void** pNetwork, const unsigned short nListenPort) | - This function is called to initialize network connection | - cond.bad() = TRUE | - Write string "Failed to initialize network." to log file
16 | int dcwSendAssociationReleaseRsp(void* association) | - This function is called to send association release response | - cond.bad() = TRUE | - Write string "Failed to send association release response." to log file
17 | int dcwAcknowledgeAssociation(void * association, const wchar_t * pszLocalAE, const wchar_t * profile) | - This function is called to send Accept /Reject association | - strncmp(asc->params->DULparams.calledAPTitle, szLocalAE, DCW_LENGTH_AE) != 0 | - Write log for print reject parameters
18 |  |  | - cond.bad() = TRUE | - Write string "Failed to reject association." to log file
19 |  |  | - pConfigFile != NULL - cond.bad() | - Write string "Failed to evaluate association." to log file
20 |  |  | - Set AE title to association params failed | - Write string "Failed to set AE title to association params." to log file
21 |  |  | - Acknowledge association failed | - Write string "Failed to acknowledge association" to log file
22 |  |  | - None | - Write log for parameters log (info log)
23 | int dcwWaitForAssociation(void* network, void** pAssociation, const int nTimeout) | - This function is called to process association waiting | - cond.bad() = TRUE | - Write string "Failed to receive association." to log file - Write log for parameters (info log)
2-1-4 DcmtkWrapCommand.cpp
\common\DcmtkWrap\DcmtkWrapCommand.cpp
No | 関数名 | 関数を呼び出す手順 | 検査条件 | 判定基準 | 備考
1 | int dcwSendNRequest(UINT nCommand, void * pAssociation, UCHAR nPresentationContextID, const char * pszSOPClassUID, const char * pszSOPInstUID, unsigned int * arrIdentifierList, const int nIdentifierListCount, void * pDataset) | - This function is called to send message N_CREATE_RQ or N_SET_RQ | - None | - Write log for dataset and request message
2 |  |  | - cond != EC_Normal | - Write string "Failed to send DIMSE message." to log file
3 | OFCondition receiveMessage(T_ASC_Association * asc, T_ASC_PresentationContextID * pid, const int nTimeout, T_DIMSE_Message * msg, DcmDataset ** rspAllTag) | - This function is called to receive message and dataset return | - nDatasetType != DIMSE_DATASET_NULL - Receive dataset failed | - Write string "Failed to recieve dataset: " to log file - Write log for received dataset and message
4 |  |  | - nDatasetType == DIMSE_DATASET_NULL | - Write log for message
5 |  |  | - rspAllTag != NULL | - Write log for received dataset and message
6 |  |  | - rspAllTag == NULL | - Write log for received dataset and message
7 | int dcwReceiveMessage(void* association, void** pDcmDataset, DIMSE_Command* command, unsigned char* pid, const int nTimeout) | - This function is called to receive command and dataset return | - cond == DIMSE_NODATAAVAILABLE | - Write string "Failed to receive DIMSE messsage." to log file
8 |  |  | - cond == DUL_PEERABORTEDASSOCIATION | - Write string "Failed to receive DIMSE messsage." to log file
9 |  |  | - cond == DUL_PEERREQUESTEDRELEASE | - Write string "Failed to receive DIMSE messsage." to log file
10 |  |  | - cond.bad() = TRUE | - Write string "Failed to receive DIMSE messsage." to log file
11 | int dcwSendCEchoRsp(void* association, unsigned char nPresentationID, int nReqMsgID, unsigned short nStatusCode) | - This function is called to send C-ECHO-RSP message | - cond.bad() = TRUE | - Write string "Failed to send C-ECHO-RSP." to the log file
12 | int dcwReceiveNEventReportRequest(void * pAssociation, void ** rspDataset, unsigned char * pPresentationContextID, const int nTimeout) | - This function is called to receive N-Event-Report request from storage commitment server | - cond == DIMSE_NODATAAVAILABLE | - Write string "Failed to receive N-EVENT-REPORT-REQUEST." to log file
13 |  |  | - cond.module() == DIMSEC_UNEXPECTEDREQUEST | - Write string "Failed to receive N-EVENT-REPORT-REQUEST." to log file
14 |  |  | - cond == DUL_PEERABORTEDASSOCIATION | - Write string "Failed to receive N-EVENT-REPORT-REQUEST." to log file
15 |  |  | - cond == DUL_PEERREQUESTEDRELEASE | - Write string "Failed to receive N-EVENT-REPORT-REQUEST." to log file
16 |  |  | - cond != EC_Normal | - Write string "Failed to receive N-EVENT-REPORT-REQUEST." to log file
17 | int dcwSendCStoreWithCallBack(void* association, void* dcmDataset, unsigned long nTag, void* pUserData, SetValueCallback callback) | - This function is called to send C-STORE-RQ with call back | - Find accepted presentation context failed | - Write string "Failed to find accepted presentation context." to log file
18 |  |  | - !obj->canWriteXfer(transXfer) | - Write string "Failed to send C-STORE-REQUEST." to log file
19 |  |  | - Send command dataset failed | - Write string "Failed to send C-STORE-REQUEST." to log file
20 |  |  | - Send pre-pixel data failed | - Write string "Failed to send C-STORE-REQUEST." to log file
21 |  |  | - Send the first frame of data failed | - Write string "Failed to send C-STORE-REQUEST." to log file
22 |  |  | - isLast == FALSE - nTmpBufsize != 0 - Send second frame of data failed | - Write string "Failed to send C-STORE-REQUEST." to log file
23 |  |  | - remainElem = TRUE - i < obj->card() | - Write log for tag value
24 |  |  | - remainElem = TRUE | - Write log for dataset format
25 | int dcwSendNEventReportRSP(void* association, unsigned char nPid, int nReqMsgID, unsigned short nStatusCode) | - This function is called to transmitting the N-EVENT-REPORT-RSP | - None | - Write log for presentation context ID
26 |  |  | - cond.bad() = TRUE | - Write string "Failed to send N-EVENT-REPORT-RSP." to log file
27 | int dcwSendCStore(void* association, void* dcmDataset) | - This function is called to send C-STORE-RQ message | - None | - Write log for dataset format
28 |  |  | - cond.bad() = TRUE | - Write string "Failed to send C-STORE-RQ." to log file
29 | int dcwSendCFindRQ(void* association, UCHAR nPresentationContextID, const char* pszSOPClassUID, void* dcmDataset) | - This function is called to send C-FIND-RQ message | - None | - Write log for dataset format and presetation context ID
30 |  |  | - cond.bad() = TRUE | - Write string "Fail to send C-FIND-RQ." to log file
31 | int dcwSendCCancelRQ(void* association, UCHAR nPresentationContextID, void* dcmDataset) | - This function is called to send C-CANCEL-RQ message | - None | - Write log for dataset format and presetation context ID
32 |  |  | - cond.bad() = TRUE | - Write string "Fail to send C-CANCEL-RQ." to log file
33 | int dcwSendCEcho(AscParamsInfo_t Params, const int nTimeout) | - This function is called to send C-ECHO command to DICOM server to check connection | - nRet != dcwERR_OK | - return nRet
34 |  |  | - None | - Write log for parameters - Write string "Sent C-ECHO-RQ." to log file
35 |  |  | - Init network failed | - Write string "Failed to initialize network." to log file
36 |  |  | - Create association failed | - Write string "Failed to association request." to log file
37 |  |  | - Send C-EHO to server failed | - Write string "Failed to send C-ECHO-RQ." to log file
38 | int dcwAddPresentationContext2AssociationConfig(const wchar_t* PresentationListKey, const wchar_t* abstractSyntaxUID, const wchar_t* transferSyntaxKey) | - This function is called to adds the given abstract syntax UID and transfer syntax list key to the list of presentation contexts | - CondRet.bad() = TRUE | - Write string "Failed to add presentation context config." to log file

## 結合テスト
2-2 結合テスト
なし

## ソフトウェアコードチェック
ソフトウェアコードチェック
次の項目に関してソフトウェアコードチェックを行う
| (A) | ソフトウェアコードが、ソフトウェアユニットの詳細設計で文書化されている インターフェースと矛盾していないか | （○：矛盾していない ×：矛盾している）
| (B) | ソフトウェアコードが、コーディング標準に従っているか | （○：従っている ×：従っていない）
| (C) | ソフトウェアコードが、実装で以下問題がないこと | （○：問題ない ×：問題あり）
・型変換によるデータ落ちがないこと
・未定義エリアの参照がないこと
| (D) | ソフトウェアユニットの詳細設計で文書化されている処理内容に対し、不足または不要な処理が無いか
| (E) | ソフトウェアユニットの詳細設計で文書化されているモジュール以外に修正を行っていないか
(モジュールごとではなく、ファイル単位で実施する。)
|  | ファイル | ユニット | （A) | （B) | （C) | （D） | （E）
|  | MntLogDlg.cpp | void OnInitDialog ()
|  | MntLogDlg.cpp | BOOL CheckDICOMLogExist()
|  | MntLogDlg.cpp | int ProcessDICOMLogFile(TCHAR * DrivePath)
|  | MntLogDlg.cpp | int CopyDICOMLogToMedia(TCHAR *szSrcFile, TCHAR *szDestFile)
|  | MntLogDlg.cpp | void OnBtnCopyToMedia1()
|  | MntLogDlg.cpp | int AddListItemFolder(TCHAR *Path)
|  | MntLogDlg.cpp | int CopyToMedia(TCHAR * DrivePath)
|  | MntTransDlg.cpp | CMntTransDlg(CWnd * pParent)
|  | MntTransDlg.cpp | BOOL OnInitDialog()
|  | MntTransDlg.cpp | void TransSCheck1()
|  | MntTransDlg.cpp | void TransSCheck2()
|  | MntTransDlg.cpp | void TransSCheck3()
|  | MntTransDlg.cpp | void TransSCheck4()
|  | MntTransDlg.cpp | void TransMCheck1()
|  | MntTransDlg.cpp | void TransMCheck2()
|  | MntTransDlg.cpp | void TransMCheck3()
|  | MntTransDlg.cpp | void TransMCheck4()
|  | MntTransDlg.cpp | void GetTransferSyntax()
|  | MntTransDlg.cpp | int GetTransferSyntaxNum(TCHAR* szAppName, TCHAR* szConfigFile)
|  | MntTransDlg.cpp | void SetTransferSyntax()
|  | DcmtkWrapAssociation.cpp | int dcwInitializeAscParam(void ** pNetwork, void ** pAscParam, const wchar_t * pszLocalAE, const wchar_t * pszLocalHost, const wchar_t * pszRemoteAE, const wchar_t * pszRemoteHost, const int nRemotePortNum, const int nTimeout)
|  | DcmtkWrapAssociation.cpp | int dcwOpenAssociation(void * network, void * assocParam, void ** pAssociation, const wchar_t * profile)
|  | DcmtkWrapAssociation.cpp | int dcwCloseAssociation(void** pAscociation)
|  | DcmtkWrapAssociation.cpp | int dcwAbortAssociation(void** pAssociation)
|  | DcmtkWrapAssociation.cpp | int dcwInitializeNetwork(void** pNetwork, const unsigned short nListenPort)
|  | DcmtkWrapAssociation.cpp | int dcwSendAssociationReleaseRsp(void* association)
|  | DcmtkWrapAssociation.cpp | int dcwAcknowledgeAssociation(void * association, const wchar_t * pszLocalAE, const wchar_t * profile)
|  | DcmtkWrapAssociation.cpp | int dcwWaitForAssociation(void* network, void** pAssociation, const int nTimeout)
|  | DcmtkWrapCommand.cpp | int dcwSendNRequest(UINT nCommand, void * pAssociation, UCHAR nPresentationContextID, const char * pszSOPClassUID, const char * pszSOPInstUID, unsigned int * arrIdentifierList, const int nIdentifierListCount, void * pDataset)
|  | DcmtkWrapCommand.cpp | OFCondition receiveMessage(T_ASC_Association * asc, T_ASC_PresentationContextID * pid, const int nTimeout, T_DIMSE_Message * msg, DcmDataset ** rspAllTag)
|  | DcmtkWrapCommand.cpp | int dcwReceiveMessage(void* association, void** pDcmDataset, DIMSE_Command* command, unsigned char* pid, const int nTimeout)
|  | DcmtkWrapCommand.cpp | int dcwSendCEchoRsp(void* association, unsigned char nPresentationID, int nReqMsgID, unsigned short nStatusCode)
|  | DcmtkWrapCommand.cpp | int dcwReceiveNEventReportRequest(void * pAssociation, void ** rspDataset, unsigned char * pPresentationContextID, const int nTimeout)
|  | DcmtkWrapCommand.cpp | int dcwSendCStoreWithCallBack(void* association, void* dcmDataset, unsigned long nTag, void* pUserData, SetValueCallback callback)
|  | DcmtkWrapCommand.cpp | int dcwSendNEventReportRSP(void* association, unsigned char nPid, int nReqMsgID, unsigned short nStatusCode)
|  | DcmtkWrapCommand.cpp | int dcwSendCStore(void* association, void* dcmDataset)
|  | DcmtkWrapCommand.cpp | int dcwSendCFindRQ(void* association, UCHAR nPresentationContextID, const char* pszSOPClassUID, void* dcmDataset)
|  | DcmtkWrapCommand.cpp | int dcwSendCCancelRQ(void* association, UCHAR nPresentationContextID, void* dcmDataset)
|  | DcmtkWrapCommand.cpp | int dcwSendCEcho(AscParamsInfo_t Params, const int nTimeout)
|  | DcmtkWrapCommand.cpp | int dcwAddPresentationContext2AssociationConfig(const wchar_t* PresentationListKey, const wchar_t* abstractSyntaxUID, const wchar_t* transferSyntaxKey)
上記ファイル以外の修正
36